Autoclave Steam Sterilization Validation Checklist

1. Document control and scope

☐ Validation plan/protocol approved

☐ Autoclave ID, location, manufacturer, model, and serial number recorded

☐ Intended use clearly defined

☐ Load types covered by the validation are listed

☐ Cycle types identified

  - Pre-vacuum / porous load

  - Gravity cycle

  - SIP cycle

  - Liquid cycle

☐ Acceptance criteria defined

☐ Deviation, change control, and CAPA procedures referenced

☐ Responsibilities assigned for Engineering, Validation, Microbiology, QA, and Operations

A defensible validation package should start with approved written procedures, defined responsibilities, and approved process controls; FDA also expects written production/process procedures and documented deviations.

2. User requirements and process definition

☐ Required sterilization objective defined

☐ Required cycle parameters defined

  - Setpoint temperature

  - Exposure time

  - Pressure range

  - Come-up / equilibration expectations

  - Drying requirements

☐ Worst-case loads identified and justified

☐ Routine load configurations documented

☐ Acceptance for wet loads defined

☐ Cooling and post-cycle contamination controls defined

☐ Product/equipment status labeling strategy defined for sterilized vs non-sterilized material

EU/WHO guidance expects the whole load to reach the required temperature before sterilizing time starts, requires post-cycle protection from recontamination, and requires a clear way to distinguish sterilized from non-sterilized items.

3. Installation Qualification (IQ)

☐ Equipment installed per approved drawings/specifications

☐ Utilities connected and identified

  - Steam / clean steam

  - Water

  - Compressed air or instrument air, if applicable

  - Power

  - Drain

☐ P&ID / wiring / loop drawings available

☐ Chamber materials of construction verified

☐ Door type and safety interlocks verified

☐ Gaskets and seals verified

☐ Instrument list complete

☐ Control system / HMI / recorder installed as specified

☐ Alarms verified

☐ Printer/electronic batch record configuration verified

☐ Software / firmware versions documented

☐ Preventive maintenance requirements identified

☐ Spare parts / critical components identified

FDA’s lifecycle validation approach places equipment qualification within Stage 2 process qualification, and autoclaves are direct-impact systems whose qualification is expected to be documented.

4. Utilities and supporting systems

☐ Steam source identified and approved for intended use

☐ Steam quality requirements defined, where applicable

☐ Steam traps and drains identified

☐ Chamber drain location verified

☐ Vacuum system identified and qualified if used

☐ Cooling media controls defined if used

☐ Vent filters identified/integrity approach defined if applicable

☐ Environmental classification/background requirements defined for loading and unloading areas, where relevant

Moist-heat performance depends on effective air removal, drainage, temperature monitoring at key points such as the chamber drain for pre-vacuum cycles, and protection from recontamination during cooling.

5. Instrumentation, calibration, and data integrity

☐ All critical instruments are identified

☐ Calibration status current for:

  - Chamber temperature probes

  - Load probes

  - Pressure instruments

  - Timers

  - Independent data loggers / thermocouples

☐ Calibration traceability documented

☐ Recorder accuracy verified

☐ Independent monitoring probe plan defined

☐ Time synchronization between autoclave recorder and validation system confirmed

☐ Electronic records / audit trail / backup controls verified, if applicable

EU/WHO guidance specifically calls for verification of control and recording probe locations using an independent monitoring probe during validation, and FDA expects qualified computerized/electronic systems where used.

6. Cycle development and load definition

☐ Each load family is defined and justified

☐ Load diagrams prepared

☐ Minimum and maximum load sizes defined

☐ Worst-case challenge loads justified

☐ Orientation, wrapping, and packaging configuration defined

☐ Drainability and steam penetration considerations assessed

☐ Loads confirmed to support effective air removal

☐ Dryness expectations defined after cycle completion

EU Annex 1 says items should be packaged so air can be removed and steam can penetrate, loads should support effective air removal and be free-draining, and loaded items should be dry when removed unless the cycle/load type justifies otherwise.

7. Operational Qualification (OQ)

☐ Empty chamber heat distribution study completed

☐ Sensor mapping plan documented

☐ Chamber cold spots / slow-to-heat locations identified

☐ Reproducibility demonstrated across repeated runs

☐ Cycle parameters challenged at approved operating ranges

☐ Alarm and abort functions challenged

☐ Door interlocks tested

☐ Power failure / restart behavior evaluated

☐ Vacuum and pressure phases verified

☐ Exposure timing logic verified

☐ Drying phase verified

☐ Chamber drain temperature monitored for pre-vacuum cycles

☐ Equilibration time calculated

☐ Pressure-temperature correlation verified

For porous cycles, EU/WHO guidance expects validation to include equilibration time, exposure time, pressure-temperature correlation, and min/max temperature range during exposure, with critical parameters under defined limits and tolerances.

8. Performance Qualification (PQ)

☐ PQ executed using routine and worst-case loads

☐ At least repeated successful runs performed per approved protocol

☐ Thermometric studies performed with load probes in worst-case positions

☐ Slow-to-heat locations confirmed

☐ Penetration to required conditions demonstrated throughout the load

☐ Exposure start confirmed only after the load reached required temperature

☐ Dryness of load assessed after cycle

☐ Packaging integrity checked after cycle

☐ Damage/distortion to load assessed

☐ Reproducibility demonstrated across repeated PQ runs

WHO states all parts of each load type should be validated by physical measurements and, where appropriate, biological indicators; EU/WHO also require the whole load to reach required temperature before sterilizing time begins.

9. Biological indicators (BI) and chemical indicators (CI)

☐ BI use is defined in the protocol

☐ BI locations represent worst-case/slow-to-heat positions

☐ BI lot numbers recorded

☐ Supplier qualification documented

☐ Storage and handling conditions controlled

☐ Positive controls included when required by procedure

☐ BI results reviewed together with physical data

☐ CIs/autoclave tape used only as supplementary evidence where appropriate

☐ Indicators are not used to override failed critical parameters

EU/WHO guidance is explicit that BI suppliers should be qualified, BI storage/handling should be controlled, and BI results alone should not override other critical parameters or process design elements; process indicators only show passage through a cycle, not sterility or SAL achievement.

10. Pre-vacuum cycle controls

☐ Leak test frequency defined

☐ Leak tests performed and accepted

☐ Air removal test method defined

☐ Daily air removal test performed for applicable pre-vacuum porous-load autoclaves

☐ Acceptance criteria for air removal established

☐ Failed leak test / air removal test response defined

EU Annex 1 states leak tests should be performed periodically, normally weekly, when a vacuum phase is part of the cycle or post-sterilization pressure drops below ambient, and it calls for an air removal test cycle normally performed daily for applicable autoclaves.

11. Routine cycle acceptance criteria

☐ Approved routine cycle printout/record reviewed for each run

☐ Time, temperature, and pressure recorded

☐ Critical limits and tolerances defined

☐ Acceptance of chamber drain temperature criteria defined for applicable cycles

☐ Load probe acceptance criteria defined where used

☐ Load dryness inspection required

☐ Package integrity inspection required

☐ Exception handling procedure defined for failed cycles

☐ Sterilization record retention requirements defined

EU/WHO guidance requires time, temperature, and pressure recording for porous cycles, inspection of items on removal, and availability of sterilization records for each run.

12. Cleaning, maintenance, and status control

☐ Cleaning procedure approved

☐ Maintenance procedure approved

☐ Cleaning and maintenance schedules established

☐ Pre-use inspection defined

☐ Chamber cleanliness inspection performed before qualification runs

☐ Equipment protected from contamination before use

☐ Maintenance/cleaning records available

☐ Status labeling in place

FDA 21 CFR 211.67 requires written cleaning and maintenance procedures, assignment of responsibility, detailed methods, protection of clean equipment from contamination, pre-use cleanliness inspection, and records of maintenance/cleaning/sanitizing/inspection.

13. Deviations and investigations

☐ All qualification deviations documented

☐ Failed probes / missing data investigated

☐ Wet loads investigated

☐ Packaging damage investigated

☐ Alarm events investigated

☐ BI/CI anomalies investigated

☐ Impact on prior or concurrent runs assessed

☐ QA disposition documented

WHO guidance requires non-conformities and deviations to be adequately investigated before release/certification, including impact assessment and justification of scope.

14. Requalification and continued verification

☐ Periodic requalification frequency defined

☐ Trigger-based requalification criteria defined:

  - Major repair

  - Relocation

  - Control system/software change

  - Utility change

  - New load family

  - Repeated cycle failures or trends

☐ Routine trend review defined

☐ Ongoing review of critical process data defined

☐ Change control linkage established

FDA’s lifecycle process validation model expects continued process verification and trend review to maintain the process in a state of control throughout its lifecycle.

15. Final validation package

☐ Approved protocol

☐ IQ/OQ/PQ records

☐ Calibration certificates

☐ Load diagrams

☐ Raw thermometric data

☐ BI/CI records

☐ Leak test and air removal test records

☐ Deviations and investigations

☐ Summary report with conclusion

☐ QA approval for validated state

☐ Approved routine SOPs linked to the validated cycle(s)

Optional add-on for terminally sterilized product loads

☐ F₀ strategy defined where applicable

☐ Heat penetration studies completed in product/load worst-case locations

☐ Container-closure integrity considerations addressed

☐ Cooling phase controls assessed for contamination risk

☐ Loading pattern shown not to distort or damage containers

☐ Parametric release strategy defined only if authorized and justified